## Slide 1
Cable Loss Calibration SOP
*Cycle: 1time/week

## Slide 2
Step 1: Run IQpathloss 1.8.1 program.

## Slide 3
Step 2: IQpathloss windows is appeared as shown below, just follow step by step:
►Select IQxel device(1) → Set IP (2) → then Connect(3).
►Set Number of chains(4) → Select port routing(5)(6) → Click Start(7)
→ connect RF1A to RF2A by RF cable then click OK(8).
1
2
3
4
5
6
7
8

## Slide 4
Step 3: Waiting for IQpathloss measuring, then click OK.

## Slide 5
Step 4: IQpathloss inform Start on Chain 1, click OK(1) → Connect the RF cable to Chain1
as the below pictures → click OK(2)
RF1A IQxel
Chain1
Devider input
Shelding box
10dB atenuatator
1
2

## Slide 6
Step 5: Waiting for Chain1 measurement results, click OK(1)
► Repeat step 5 until IQpathloss inform Start on Chain 5, insert RF cable to Chain2
► Repeat step 5 until IQpathloss inform Start on Chain 9, insert RF cable to Chain3
► Repeat step 5 until IQpathloss inform Start on Chain 13, insert RF cable to Chain4
► Repeat step 5 until IQpathloss inform Cable loss measurement done., click OK(2)
Chain1
Chain2
Chain3
Chain4
Power divider
1
2

## Slide 7
Step 6: Run CableLossTool_2.0.0.1 program.

## Slide 8
Step 7: Click select(1) → Link to path_loss file(2) → Click Open(3).
1
2
3

## Slide 9
Step 8: Cable loss result PASS/FAIL will be shown as the below:
►On the case FAIL, check again the connections then proceed cable loss again.
►On the case PASS, move to Step 9.
PASS
FAIL

## Slide 10
Step 9: Set Shift value(1) → Click Shift CSV(2).
1
2

## Slide 11
Step 10: Link path_loss-Shift-yyyymmdd file to the ‘C:\LitePoint\Iqpathloss\1.8.1’ (1)→ Save(2)
path_loss-Shift-yyyymmdd csv file is saved. Cable Loss Calibraton is DONE.
1
2

## Slide 12: Thank you!